# 510(k) Clearance Summary — K243178

Sample document. K-number is illustrative.

# Clearance

Submission number: K243178

Decision date: August 2024

Decision: Substantially Equivalent (SE) — cleared

Submission type: Traditional 510(k)

Regulation number: 21 CFR 892.2050 (Picture archiving and communications system)

Product code: QIH — Radiological computer-assisted diagnostic software for lesions suspicious for cancer (used by analogy for AI-CADe/CADx imaging software; FDA accepted)

Class: II

Submitter: HealthCo, Inc.

# Device

EchoPilot is a software-only medical device (SaMD) that analyzes transthoracic echocardiogram studies and produces a probability score that a study contains imaging features consistent with cardiac amyloidosis. EchoPilot is intended for use as an adjunctive tool by qualified cardiologists; it does not replace independent clinical judgment.

# Predicate

Ultromics EchoGo Pro (K233412, cleared 2023) was selected as the primary predicate. EchoGo Pro is also a Class II SaMD analyzing echocardiograms for structural heart disease features. The substantial-equivalence rationale rested on:

Same intended use category (adjunctive AI analysis of echo studies)

Same target user population (cardiologists)

Same fundamental technological characteristics (deep-learning analysis of DICOM echo)

Comparable performance characteristics demonstrated through bench testing

A reference device — Us2.ai Us2.v1 (K221XXX) — was cited for additional support on the validation methodology.

# Indications for Use

EchoPilot is a software-only medical device intended to assist trained cardiologists in the detection of imaging features consistent with cardiac amyloidosis from transthoracic echocardiogram studies acquired on commercially available ultrasound systems. EchoPilot outputs a probability score and feature-attribution overlay. EchoPilot is adjunctive to, and not a replacement for, the clinical judgment of qualified physicians. EchoPilot is intended for use in adult patients (18+).

Off-label exclusions explicitly stated: not for pediatric use, not for stress-echo or transesophageal studies, not for primary diagnosis (adjunctive only).

# Performance Testing Submitted

The 510(k) included:

Standalone analytical validation — bench testing on a sequestered test set of 1,840 echo studies, demonstrating sensitivity 91.4% and specificity 87.9% against an expert-adjudicated reference standard

Software verification & validation — IEC 62304 Class B documentation

Cybersecurity — premarket cybersecurity package per FDA 2023 guidance

Usability / human factors — IEC 62366 summative validation with 15 cardiologist participants

Bench comparison to predicate — head-to-head on a common test set

# FDA Interactions

The submission cycle included one Additional Information request (October 2023) addressing demographic subgroup analysis and labeling clarifications. HealthCo's response was accepted without further AI requests.

# Post-Market Commitments

None. No post-market study (522 order) was required. HealthCo voluntarily maintains a post-market surveillance dashboard tracking field performance across pilot sites; this is internal and not an FDA commitment.

# Labeling

The cleared labeling is reproduced in full at Tier 3 (labeling-IFU-final.pdf). Headline excerpts are above.