# Regulatory Pathway Roadmap

Sample document. Forward-looking; subject to FDA timing.

This roadmap lays out HealthCo's regulatory plan through 2028, anchored to the commercial roadmap in 06-Market-Competitive/.

# Current State (March 2026)

EchoPilot core (amyloidosis detection) — 510(k) cleared K243178 (August 2024). Locked algorithm, no PCCP. Fully commercial.

Auto-strain feature — De Novo accepted Q1 2026; substantive review underway. Expected clearance: late FY26 (Q4), plausible push to early FY27.

QMS — Certified to ISO 13485:2016 (April 2024, Lloyd's Register); first surveillance audit passed November 2025 with no nonconformances.

MDR / EU-MDR — CE mark: not pursued (deferred; see below).

# FY26 Plan

| Initiative | Pathway | Target Decision |
| --- | --- | --- |
| Auto-strain De Novo | De Novo with PCCP | Q4 FY26 |
| Cybersecurity update (post-2024 FDA guidance refresh) | Letter-to-file | Q3 FY26 |
| Pediatric exclusion language refinement | Letter-to-file | Q2 FY26 |

# FY27 Plan

| Initiative | Pathway | Target Decision |
| --- | --- | --- |
| Cardio-oncology indication (chemotherapy-induced cardiotoxicity surveillance) | New 510(k) | Q4 FY27 |
| Hypertrophic cardiomyopathy detection | 510(k) supplement | Q3 FY27 |
| EU-MDR (Class IIa via notified body) — gated on Series A close | CE mark / MDR | FY28 |

The cardio-oncology and HCM indications were the subject of the October 2025 Q-Sub (Q251014). FDA's feedback was constructive; both pathways are operationalized in the FY26-27 budget.

# FY28 Plan

| Initiative | Pathway | Target Decision |
| --- | --- | --- |
| Pediatric indication (initial) | New 510(k) | Q4 FY28 |
| Stress-echo extension | 510(k) supplement | Q3 FY28 |
| Japan PMDA submission | PMDA Class II | FY28 |

# Quality Management System

ISO 13485:2016 certified (Lloyd's Register), April 2024

IEC 62304 SDLC compliance — Class B

IEC 62366 usability engineering processes

ISO 14971 risk management — current risk file maintained

ISO/IEC 27001 information security — certification targeted Q4 FY26 (in scope of current external audit)

# Predetermined Change Control Plan (PCCP)

The De Novo submission includes a PCCP with three categories of pre-authorized changes:

Performance-equivalent data refresh — adding new training data within prespecified composition bounds; performance bounds prespecified

Vendor-calibration table updates — adding new echo-machine vendor/version characterizations

Confidence-threshold tuning — adjusting the confidence-gating threshold within validated ranges

The PCCP does not authorize architectural changes, indication changes, or moves to fully autonomous (non-adjunctive) operation. Those continue to require new submissions.

# Internal Governance

VP Regulatory & Quality — owns roadmap and submissions

Regulatory Steering Committee — monthly; CEO, COO, CTO, VP Reg/Quality

Quality Management Review — quarterly per ISO 13485 §5.6

External advisor — Hogan Lovells partner attends quarterly RSC meetings