# Software Development Process

Sample document. Describes HealthCo's SDLC and software-quality framework as it applies to EchoPilot.

EchoPilot is a Class II medical device classified as SaMD. HealthCo's software development process is structured to satisfy:

IEC 62304:2006/AMD1:2015 — software lifecycle for medical device software (Class B)

ISO 13485:2016 — quality management for medical devices

ISO 14971:2019 — risk management

IEC 62366-1:2015 — usability engineering

FDA 2023 Premarket Cybersecurity Guidance

FDA Good Machine Learning Practice (GMLP) Guiding Principles

# SDLC Overview

Development follows a documented lifecycle with controlled phases and gated exits:

Concept and risk classification — feature scoped against intended use; safety classification confirmed (Class B)

Software requirements — requirements traced from clinical user needs through verification

Architectural design — module-level design, interface contracts, security architecture

Detailed design and implementation — coding standards, peer review on every PR

Verification — unit, integration, system, and analytical performance testing

Validation — usability, clinical performance, real-world workflow validation

Release — controlled release with signed artifacts, version-locked dependencies, and a release record

Post-market surveillance — field-monitoring dashboard, periodic safety updates

Every requirement, design element, test, and risk control is traced through a Requirements Traceability Matrix (RTM) maintained in Jama Connect.

# Version Control

Source code — GitHub Enterprise Cloud, organization-scoped, branch protection on main

Mandatory reviewers — every PR requires (a) code review by a non-author engineer and (b) for safety-relevant code, additional review by a designated regulatory-aware engineer

Signed commits — required on main and release branches

Immutable releases — release artifacts (model weights, container images) are content-addressed and signed with Sigstore; the release record references the exact artifact digest

# Change Management

Changes are categorized at design-time:

Maintenance change (e.g., dependency patch) — letter-to-file under cybersecurity update workflow if security-relevant

Within-PCCP change (post De Novo clearance) — handled per the Predetermined Change Control Plan

Outside-PCCP change — triggers either a 510(k) supplement or a new submission

A cross-functional Change Control Board (CTO, VP Reg/Quality, VP Clinical) reviews every change with potential safety or performance impact before merge.

# Verification and Validation

Unit tests — coverage gates enforced (>85% statement coverage on safety-classified modules)

Integration tests — automated against simulated DICOM workflows

System tests — end-to-end on representative echo studies, executed pre-release

Analytical performance test — sequestered test set, run on every model release; performance versus locked acceptance criteria

Cybersecurity test — SAST in CI, DAST quarterly, third-party penetration test annually (most recent: November 2025, by NCC Group; remediation to clean within agreed SLAs)

# Risk Management

ISO 14971 risk file is a living document. Every identified hazard is decomposed to hazardous situations and harms, with risk controls, residual-risk justification, and verification of each control. The risk file is reviewed at every release and at the quarterly Quality Management Review.

# Coding Standards and Tooling

Languages: Python (model serving and orchestration), TypeScript (clinician-facing UI), CUDA (GPU kernels for inference)

Static analysis: Ruff, mypy, eslint, semgrep

Dependencies: SBOM generated per release (Syft); vulnerability scanning (Grype) in CI; license-compliance check

Reproducibility: training runs containerized; data lineage tracked end-to-end via Pachyderm-style hashing

# Cybersecurity

Threat modeling per release (STRIDE-based)

Secure-by-default architecture: zero plaintext PHI at rest, least-privilege IAM, mutual TLS for service-to-service

Annual third-party penetration test (NCC Group)

Vulnerability disclosure program (security@healthco.example)

Incident response runbook tested twice yearly

# Audit History

ISO 13485 certification audit, April 2024 — passed

ISO 13485 surveillance audit, November 2025 — no nonconformances

Internal QMS audit, monthly per documented schedule